## Slide 1: DOAC-Stop™ Research Use Only. Not for use in diagnostic procedures.
Jennifer J. Kiblinger
Scientific Director

## Slide 2: What is DOAC-Stop™?
It’s a sample pre-treatment
DOAC-Stop™ is the first general agent to neutralize DOACs (Direct Oral Anticoagulants) from plasma, for simplified testing
DOAC-Stop™ efficiently eliminates DOACs (Direct Oral Anticoagulants) including dabigatran, apixaban, rivaroxaban and edoxaban, from test plasmas with minimal effect on currently-known clotting variables

## Slide 3: How does it work?
  Adsorbent product (proprietary composition based on activated charcoal) that extracts DOACs from plasmas
  Simple procedure:
    add one mini-tablet or 20µL DOAC-Stop Liquid per 1ml test plasma (volume is not critical) or one micro-tablet per 0.5ml test plasma
    Mix for 5 minutes and centrifuge down the adsorbent.
  The resulting plasma is then depleted of DOACs
  Yields similar APTT, dRVVT, PT(INR) and other results as if DOACs had not been present.

## Slide 4: Tips
Mix with test plasmas for 5 minutes after the mini or micro-tablet has fully dispersed, NOT from the time of its addition
Shake DOAC-Stop Liquid vigorously before use
A longer incubation time could be used if not fully dispersed within 5 minutes, although studies show still effective.

## Slide 5: DOAC-Stop Applications
Plasmas treated with DOAC-Stop™ can be used for reliable routine assessment of known blood coagulation parameters in research settings.
The degree of shortening induced by DOAC-Stop in a clotting time test depends on the concentration and type of DOAC as well as the on the underlying plasma itself.

## Slide 6: DOAC-Stop Applications
Clotting time results with DOAC sensitive tests such as dRVVT and APTT tests, may be expressed as ratios of test result before to that obtained after DOAC-Stop™ treatment.
This ratio ranges upward from 1.0 and correlates positively with DOAC concentration depending on which DOAC and test is used.
This “Correction Ratio” where the clotting test result before sample treatment by DOAC-Stop™ is divided by the clotting test results after treatment by DOAC-Stop™, may be useful with high phospholipid dRVVT reagents

## Slide 7: How much does it remove?
  DOAC-Stop™ mini tabs absorb up to an estimated 2,000ng/ml of any DOAC in less than 10 minutes.
  Leaves no residual effect

## Slide 8: What About Heparin and Warfarin?
There is negligible interference with vitamin K antagonist or heparinoid anticoagulants.
Works On | May Work | Doesn’t Work
Dabigatran | Hirudin* | Heparin
Apixaban | Bivalirudin* | Warfarin
Rivaroxaban | Lepirudin* | 
Edoxaban | Argatroban* | 
*Bivalirudin is similar to hirudin and lepirudin and has shown to be extracted by DOAC-Stop from plasmas, although more slowly than DOACs, being of higher MW.

## Slide 9: What about other proteins?
Specificity for neutralizing DOACs is at least partly due to molecular size.
Practically all proteins, clotting factors and inhibitors of physiological origin affecting coagulation tests have molecular size above approximately 30,000 kDa.
The adsorbing agent in DOAC-Stop™ has a pore capacity of approximately 5000
Therefore, activated charcoal-based product extracts DOACs efficiently with no effect on heparin-type anticoagulants.
Each lab must validate the assay to ensure no DOAC-Stop™ binding of other proteins that could affect other assays

## Slide 10: Packaging
Each mini-tablet or 20µL DOAC-Stop Liquid contains enough DOAC-Stop™ agent for processing 1ml of citrated test plasma (0.5-1.5ml is acceptable).
Each micro-tablet contains enough DOAC-Stop™ agent for processing of 0.5ml of citrated test plasma
Products are stable for three years at room temperature.
See Instructions for Use for specific storage conditions.
DOAC-Stop™ mini-tabs | X9904-50 | 50 mini-tabs
DOAC-Stop™ mini-tabs | X9904-100 | 100 mini-tabs
DOAC-Stop™ micro-tabs | X9904µ-50 | 50 mini-tabs
DOAC-Stop™ micro-tabs | X9904µ-100 | 100 micro-tabs
DOAC-Stop™ Liquid | X9905-100 | 1 x 2mL vial

## Slide 11: InhibiRef Tubes
InhibiRef Tubes provide a pre-determined quantity of a coagulation
  inhibitor such as a DOAC or heparin
They allow laboratories to prepare inhibitor samples at any concentration
  in any biological fluid for use in research, teaching or other purposes.
Available as rapidly-dissolving dry film in a vial.
The inhibitors dissolve within 10 minutes into 1ml of fluids such as blood, plasma or urine
The compound is therefore transferred into samples.
InhibiRef in vials makes it easier to mix and store the sample.
Plasmas treated using InhibiRef may be used as positive controls in clotting, chromogenic or other assays.
Dilutions can be prepared in pooled normal plasma, blood or other fluids to set up calibration curves.​
All inhibitor types are color-coded for recognition. These customized inhibitors can be manufactured on demand based on your laboratory needs.
Samples prepared using InhibiRef can be used for checking the efficacy of DOAC-Stop™.
Samples prepared with DOACs from InhibiRef should show restoration of original test results after treatment with DOAC-Stop™. Samples prepared with Heparin Ref should not be modified by DOAC-Stop™.​​

## Slide 12: In Conclusion
Removal by extraction of DOACs by DOAC-Stop™ allows research labs to perform clotting assays that are affected by the presence of DOACs
It can be used for checking if DOACs are present in research samples, giving unexpectedly prolonged clotting test or abnormal chromogenic test results
Research plasmas treated with DOAC-Stop™ can be used for valid factor assays
It can prevent false LAC positive results due to DOACs in testing for lupus anticoagulants in a plasma sample in a research setting

## Slide 13: Questions? Contact a DiaPharma Scientist Today!
ML-00-00667Rev02